## Calculation
Calculation of inspiration restrictor resistance to alter patient ventilation characteristics
Of use for ventilator sharing between patients |  | White fields are for data entry
[7 empty] | Time (s) | 1 | 2 | 3 | 4
Ventilator settings. Ventilator must in Pressure Control mode |  |  |  |  | Derived ventilator parameters |  | 0 | 0.4607 | 0.34515 | 0.23808 | 0.14507
Respiration rate (breaths / min) | 15 |  |  |  | Respiration period (s) | 4 | 0.05 | 0.49078 | 0.37567 | 0.26893 | 0.17575
Inspiration to expiration time ratio (I/E) | 0.5 |  |  |  | I to E switching time within period (s) | 1.3333 | 0.1 | 0.52008 | 0.40523 | 0.29857 | 0.20478
Pinsp (inspiration pressure above PEEP, cmH2O) | 25 |  |  |  | Max ventilator pressure (cmH2O) | 30 | 0.15 | 0.54861 | 0.43387 | 0.32703 | 0.23227
PEEP (cmH2O) | 5 | [5 empty] | 0.2 | 0.57641 | 0.46161 | 0.35438 | 0.25829
Tubing system resistance (cmH2O/(L/s) ) | 20 | [5 empty] | 0.25 | 0.60349 | 0.48848 | 0.38064 | 0.28292
Tubing system compliance (L/cmH2O) | 0.006 | [5 empty] | 0.3 | 0.62986 | 0.51452 | 0.40587 | 0.30624
[7 empty] | 0.35 | 0.65555 | 0.53973 | 0.4301 | 0.32831
Patient characteristics | A | B | C | D |  |  | 0.4 | 0.68057 | 0.56416 | 0.45338 | 0.34921
Inspiration resistance, Ri (cmH2O/(L/s) ) | 13 | 13 | 13 | 13 | Note: 13 is median for normal lung, Arnal 2018 |  | 0.45 | 0.70495 | 0.58782 | 0.47574 | 0.36899
Expiration resistance, Re (cmH2O/(L/s) ) | 13 | 13 | 13 | 13 | Note: same as inspiration often used |  | 0.5 | 0.72869 | 0.61074 | 0.49721 | 0.38771
Inspiration Compliance, Ci (L/cmH2O) | 0.054 | 0.044 | 0.034 | 0.024 | Note: 0.054 is median for normal lung, Arnal 2018 |  | 0.55 | 0.75182 | 0.63295 | 0.51784 | 0.40544
Expiration Compliance, Ce (L/cmH2O) | 0.054 | 0.044 | 0.034 | 0.024 | Note: same as inspiration often used |  | 0.6 | 0.77434 | 0.65446 | 0.53765 | 0.42222
Derived patient parameters | [6 empty] | 0.65 | 0.79628 | 0.67529 | 0.55669 | 0.43811
Patient inspiration time constant, RiCi (s) | 0.702 | 0.572 | 0.442 | 0.312 |  |  | 0.7 | 0.81765 | 0.69548 | 0.57497 | 0.45315
Patient expiration time constant, ReCe (s) | 0.702 | 0.572 | 0.442 | 0.312 |  |  | 0.75 | 0.83847 | 0.71503 | 0.59253 | 0.46738
Derived system parameters | [6 empty] | 0.8 | 0.85875 | 0.73397 | 0.6094 | 0.48086
Ventilation inspiration time constant (s) | 1.902 | 1.572 | 1.242 | 0.912 |  |  | 0.85 | 0.8785 | 0.75231 | 0.6256 | 0.49362
Ventilation expiration time constant (s) | 1.902 | 1.572 | 1.242 | 0.912 |  |  | 0.9 | 0.89774 | 0.77008 | 0.64116 | 0.5057
Steady end-inspiration lung volume (L) | 1.62 | 1.32 | 1.02 | 0.72 |  |  | 0.95 | 0.91648 | 0.7873 | 0.65611 | 0.51713
Steady end-expiration lung volume (L) | 0.27 | 0.22 | 0.17 | 0.12 |  |  | 1 | 0.93473 | 0.80398 | 0.67047 | 0.52795
Coefficient a | 0.027212 | 0.025159 | 0.022379 | 0.018437 |  |  | 1.05 | 0.95251 | 0.82013 | 0.68426 | 0.5382
Coefficient b | 0.49608 | 0.4282 | 0.3418 | 0.23177 |  |  | 1.1 | 0.96983 | 0.83578 | 0.69751 | 0.5479
Coefficient c | 0.040711 | 0.035933 | 0.030028 | 0.022711 |  |  | 1.15 | 0.9867 | 0.85094 | 0.71023 | 0.55708
Coefficient d | 0.2461 | 0.18335 | 0.11683 | 0.05372 |  |  | 1.2 | 1.0031 | 0.86562 | 0.72245 | 0.56577
Max lung volume, in addition to FRC (L) | 1.0449 | 0.90257 | 0.75274 | 0.58675 |  |  | 1.25 | 1.0191 | 0.87985 | 0.73419 | 0.574
Min lung volume, in addition to FRC (L) | 0.4607 | 0.34515 | 0.23808 | 0.14507 |  |  | 1.3 | 1.0347 | 0.89363 | 0.74547 | 0.58179
Tidal volume (L) | 0.58419 | 0.55742 | 0.51466 | 0.44167 |  |  | 1.35 | 1.0381 | 0.89538 | 0.74497 | 0.57829
[7 empty] | 1.4 | 1.0182 | 0.87423 | 0.72229 | 0.55384
Inspiration restriction to achieve a specified tidal volume | [6 empty] | 1.45 | 0.99879 | 0.85375 | 0.70049 | 0.5307
Required tidal volume (L). Must be <= value above | 0.4415 | 0.4415 | 0.4415 | 0.4415 |  |  | 1.5 | 0.97988 | 0.83391 | 0.67956 | 0.50879
New max lung volume, in addition to FRC (L) | 0.85562 | 0.76062 | 0.6699 | 0.58656 |  |  | 1.55 | 0.96146 | 0.81469 | 0.65946 | 0.48805
New min lung volume, in addition to FRC (L) | 0.41412 | 0.31912 | 0.2284 | 0.14506 |  |  | 1.6 | 0.94352 | 0.79607 | 0.64014 | 0.46841
New coefficient b | 0.63388 | 0.55889 | 0.44227 | 0.23209 |  |  | 1.65 | 0.92605 | 0.77804 | 0.62159 | 0.44983
New ventilation inspiration time constant (s) | 2.9246 | 2.2917 | 1.6343 | 0.91285 |  |  | 1.7 | 0.90903 | 0.76057 | 0.60377 | 0.43223
Inspiration restrictor needed (cmH2O/(L/s)) | 17.0437 | 14.3941 | 9.8077 | 0.028174 |  |  | 1.75 | 0.89245 | 0.74365 | 0.58666 | 0.41557
OR inspiration added compliance needed (L/cmH2O) | 0.051131 | 0.035985 | 0.019615 | 4.226e-05 |  |  | 1.8 | 0.8763 | 0.72725 | 0.57022 | 0.3998
[7 empty] | 1.85 | 0.86057 | 0.71137 | 0.55442 | 0.38488
Inspiration restriction to achieve the same tidal volume (row 32) with a higher Pinsp pressure | [6 empty] | 1.9 | 0.84524 | 0.69599 | 0.53926 | 0.37075
Higher Pinsp (insp. pressure above PEEP, cmH2O) | 30 | 30 | 30 | 30 |  |  | 1.95 | 0.83032 | 0.68109 | 0.52469 | 0.35737
New steady end-inspiration lung volume (L) | 1.89 | 1.54 | 1.19 | 0.84 |  |  | 2 | 0.81578 | 0.66665 | 0.51069 | 0.34471
New coefficient b | 0.59127 | 0.53348 | 0.45934 | 0.36443 |  |  | 2.05 | 0.80162 | 0.65267 | 0.49725 | 0.33272
New ventilation inspiration time constant (s) | 2.5374 | 2.122 | 1.7139 | 1.3209 |  |  | 2.1 | 0.78783 | 0.63913 | 0.48434 | 0.32137
Inspiration restrictor needed (cmH2O/(L/s)) | 10.5896 | 11 | 11.7972 | 13.6299 |  |  | 2.15 | 0.77439 | 0.626 | 0.47193 | 0.31063
[7 empty] | 2.2 | 0.7613 | 0.61329 | 0.46002 | 0.30046
[7 empty] | 2.25 | 0.74856 | 0.60098 | 0.44857 | 0.29083
[7 empty] | 2.3 | 0.73614 | 0.58905 | 0.43758 | 0.28172
[7 empty] | 2.35 | 0.72405 | 0.5775 | 0.42702 | 0.27309
[7 empty] | 2.4 | 0.71227 | 0.56631 | 0.41688 | 0.26492
[7 empty] | 2.45 | 0.70079 | 0.55547 | 0.40714 | 0.25719
[7 empty] | 2.5 | 0.68961 | 0.54496 | 0.39778 | 0.24987
[7 empty] | 2.55 | 0.67873 | 0.53479 | 0.3888 | 0.24294
[7 empty] | 2.6 | 0.66812 | 0.52494 | 0.38016 | 0.23638
[7 empty] | 2.65 | 0.65779 | 0.51539 | 0.37187 | 0.23018
[7 empty] | 2.7 | 0.64773 | 0.50614 | 0.3639 | 0.2243
[7 empty] | 2.75 | 0.63793 | 0.49718 | 0.35625 | 0.21873
[7 empty] | 2.8 | 0.62838 | 0.48851 | 0.3489 | 0.21347
[7 empty] | 2.85 | 0.61909 | 0.4801 | 0.34185 | 0.20848
[7 empty] | 2.9 | 0.61003 | 0.47196 | 0.33506 | 0.20376
[7 empty] | 2.95 | 0.60121 | 0.46407 | 0.32855 | 0.19929
[7 empty] | 3 | 0.59261 | 0.45643 | 0.3223 | 0.19506
[7 empty] | 3.05 | 0.58424 | 0.44903 | 0.31629 | 0.19106
[7 empty] | 3.1 | 0.57609 | 0.44186 | 0.31051 | 0.18727
[7 empty] | 3.15 | 0.56815 | 0.43491 | 0.30497 | 0.18368
[7 empty] | 3.2 | 0.56041 | 0.42818 | 0.29964 | 0.18028
[7 empty] | 3.25 | 0.55288 | 0.42167 | 0.29453 | 0.17706
[7 empty] | 3.3 | 0.54554 | 0.41535 | 0.28961 | 0.17402
[7 empty] | 3.35 | 0.53839 | 0.40924 | 0.28489 | 0.17114
[7 empty] | 3.4 | 0.53143 | 0.40331 | 0.28036 | 0.16841
[7 empty] | 3.45 | 0.52464 | 0.39757 | 0.27601 | 0.16583
[7 empty] | 3.5 | 0.51804 | 0.39202 | 0.27182 | 0.16338
[7 empty] | 3.55 | 0.5116 | 0.38663 | 0.26781 | 0.16107
[7 empty] | 3.6 | 0.50533 | 0.38141 | 0.26395 | 0.15888
[7 empty] | 3.65 | 0.49923 | 0.37636 | 0.26024 | 0.1568
[7 empty] | 3.7 | 0.49328 | 0.37147 | 0.25668 | 0.15484
[7 empty] | 3.75 | 0.48749 | 0.36672 | 0.25326 | 0.15298
[7 empty] | 3.8 | 0.48184 | 0.36213 | 0.24997 | 0.15122
[7 empty] | 3.85 | 0.47635 | 0.35768 | 0.24682 | 0.14956
[7 empty] | 3.9 | 0.47099 | 0.35337 | 0.24379 | 0.14798
[7 empty] | 3.95 | 0.46578 | 0.3492 | 0.24088 | 0.14649
[7 empty] | 4 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.05 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.1 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.15 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.2 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.25 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.3 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.35 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.4 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.45 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.5 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.55 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.6 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.65 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.7 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.75 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.8 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.85 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.9 | 0.4607 | 0.34515 | 0.23808 | 0.14507
[7 empty] | 4.95 | 0.4607 | 0.34515 | 0.23808 | 0.14507
Note on A41: Andy: Alternative modification to get vol. of Row 35. Caveat: to affect inspiration only, needs to be distal side of exp. check valve, resistance of which has not been accounted for.

## Description
Calculation of inspiration restrictor resistance to alter patient ventilation characteristics
Of use for ventilator sharing between patients
A R Plummer, University of Bath, 2nd April 2020
v2 4/4/2020
A patient can be characterised by a series resistance (R) and compliance (C). Ventilator tubing system can also be approximated by lumped linear model, i.e resistance (Rv) and compliance (Cv). When two patients are hooked up to the same ventilator, pressure control will be used, the ventilator stepping between low pressure (PEEP) and high pressure (PEEP+Pinsp). So the tidal volume (Vmax-Vmin) delivered to each patient, from minimum lung volume during expiration (Vmin), to maximum lung volume during inspiration (Vmax), can no longer be individually controlled. This sheet calculates the volumes for each patient. Also, for given desired tidal volumes, the values of additional restrictions (resistances) in the patients' inspiration lines can be calculated. (Note: This could plausibly also be done by adjusting inspiration compliances) NOTE : lung volume is taken as zero at zero gauge pressure. FRC is not included.
The linear lumped resistance-compliance network is is a highly simplified model with four parameters: linear resistance and compliance for both ventilator tubing system and patient. Different parameters can be used for inspiration and expiration phases (so 8 parameters in total). The simple RC network model, even when approximated further to a 1st order lag, seems to match much more sophisticated physics-based models well, but whether this is consistently the case with constant parameter values has yet to be determined. The benefit of the RC model is that simple calculations are possible on how patient R and C variation will effect tidal volume, and how this might be corrected by adding a restrictor valve (or possibly other modifications). One scenario could be to calculate what inspiration restriction is required to operate with an increased ventilator pressure (which may be necessitated by ventilator sharing) without increasing the patient’s tidal volume.
EXAMPLE. An example ventilator pressure control setting might be described as 30/15 cmH20, meaning Pinsp=15 cmH20, and PEEP = 15 (so ventilator switches between 30 and 15, because Pinsp is the pressure rise above PEEP). Then you get different tidal volume ranges (row 32) for the 4 patients with these pressure settings. Adding an inspiration restrictor can only reduce the tidal volume, not increase it. Putting the same value as the calculated tidal volume in the entry field for tidal volume (row 35) will give a restrictor resistance of zero, or a smaller value gives a sensible restriction. Putting a larger tidal volume gives a negative restriction (as its trying to supercharge the patient to get more volume in! ... or in the extreme it can't calculate a value at all).